## Intro&Application
PART 3: TEST REPORT FORMAT FOR TYPE EVALUATION
INTRODUCTION
This Report Format applies for any kind of protein measuring instrument for grain (independent of its technology). It presents a standardized format for the results of the various tests and examinations, outlined in Part 2 clause 7 of OIML R xxx (201x), to which a type of protein measuring instrument for grain shall be submitted with a view to its approval based this OIML Recommendation.
It is recommended that all metrology services or laboratories evaluating and/or testing types of protein measuring instrument for grain according to OIML R xxx (201x), or to national or regional regulations based on that Recommendation, use this Report Format, directly or after translation into a language other than English or French. In case of a translation, it is highly recommended to leave the structure and the numbers of the clauses unchanged: in this case most of the contents are also understandable for those who cannot read the language of the translation.
It is also recommended that this Report Format in English or in French (or in both languages) be transmitted by the country performing the tests to the relevant authorities of another country, when requested for issuing a national or regional type-approval.
In the practical application of the Report Format, a cover page shall be included by the Issuing Authority, and clauses 1 - 5 shall be included as a minimum.
APPLICABILITY OF THIS TEST REPORT FORMAT
In the framework of the OIML Certificate System for Measuring Instruments applicable to protein measuring instruments for grain in conformity with OIML R xxx (201x), use of this report format is mandatory, in French and/or in English with translation into the national languages of the countries issuing such certificates, if appropriate.
Implementation of this Report Format is informative with regard to the implementation of OIML Recommendation OIML R xxx (201x) in national regulations.
GUIDANCE FOR THE APPLICATION OF THIS TEST REPORT
Refer to Part 1 clause 2 of of OIML R xxx (201x) for definitions of terms, acronyms and symbols used.
Some abbreviations and symbols specific to a fewpages within Part 3 Test Report include:
H heat (e.g. 'Dry H 'is 'Dry heat' and 'Damp H' is 'Damp heat')
GT grain type (e.g. wheat, barley)
ID unique identifier
Ref reference (typically 'reference conditions' but occasionally 'reference PMB value')
Unom nominal test voltage
V voltage (e.g. 'Low V' is 'Low voltage' and 'High V' is High voltage)
The measurement unit is not always stated in the header row or column in the table.
The examiner is not expected to include the measurement unit with each recorded measurement result.
Percent by weight (abbreviated as' % w/w' or '%') is the measurement unit applicable for any values of the following:
- Protein content (PMB)
- MPE, error shift limits, maximum fault, etc.
- Basis moisture content (MB), and
- Actual "as is" moisture content (M)
The coloured fields should always be filled as appropriate.
Where the heading or label of a coloured field indicates "pass/ fail" or "pass/ fail/ NA", select the applicable option in the drop-down list that appears when the mouse is placed over the right side of the field.
NOTE: If it is impossible to enter the results in a computer, this Report Form may be printed and completed manually. In this case write "pass", "fail" or "NA" in the coloured fields as appropriate.
In the Examination Checklist, the optional comments fields have a different colour to the mandatory fields (i.e. fields labelled "yes/ no", "pass/ fail" response or details of validation)
The mandatory fields relating to conditions to be specified by the national responsible body (within the limits suggested in OIML R xxx) have a different colour to the fields with the result from the assessment of the submitted type and documentation (i.e. fields labelled "pass/ fail/ NA"). Select "NA" only when the requirement (or a variant of a requirement) is not adopted by the national responsible body.
In case a prescribed test or requirement is not relevant for the type of instrument to be tested or has not been adopted by the national responsible body, the reason why the test is omitted shall be clearly stated in the field “Comments”
In the test reports, the white fields with blue outline contain calculations and/or conditional formatting to highlight a suspect result.
THE EVALUATION REPORT
The format for the report is given on the following pages
To review or revise the formulae for calculated values in the file 201407 - 5CD Protein in grain - Part 3 Test Report.xlsx, select Menu Option 'Review' and then select 'Unprotect Sheet'.

## Cover-IssuingAuthority
COVER PAGE BY THE ISSUING AUTHORITY (NATIONAL RESPONSIBLE BODY)

## Contents
CONTENTS OF PART 3: TEST REPORT FORMAT | Location
Introduction | Intro&Application
Applicability of this Test Report Format | Intro&Application
Guidance for the application of this Test Report | Intro&Application
The evaluation report | Intro&Application
1 Authority responsible for this Report | Authority&Synopsis
2 Synopsis of the results of the examination and tests | Authority&Synopsis
3 Summary of the results of the examination and tests
3.1 Examinations | Summary
3.2 Performance tests | Summary
4 General information about the application
4.1 Manufacturer | ApplicationGenInfo
4.2 Applicant | ApplicationGenInfo
4.3 Testing laboratories involved in the tests | ApplicationGenInfo
4.4 General information concerning the type | ApplicationGenInfo
4.5 Accessories supplied by the applicant | ApplicationGenInfo
4.6 Selection of sample instruments | ApplicationGenInfo
4.7 Adjustments and modifications | ApplicationGenInfo
4.8 Results of previous tests that were taken into account | ApplicationGenInfo
4.9 Additional information on the submitted type | ApplicationGenInfo
4.10 Documentation supplied by applicant | ApplicationGenInfo
4.11 Test equipment and grain used in type evaluation | ApplicationGenInfo
5 Examination details
5.1 Checklist – specifications to be confirmed by the national authority | SpecConfirm
5.2 Examination checklist – metrological and technical requirements | ExamDetails
5.3 Examination checklist – software-controlled devices and security | ExamDetails
5.4 Examination checklist – documentation requirements | ExamDetails
6 Type evaluation tests
6.1 Tests for time-related effects
6.1.1 Instrument warm-up time | WarmUp
6.1.2 Instrument drift and instability | DriftInstability
6.2 Tests for influence factor variations within the rated operating ranges
6.2.1 Instrument levelling | Levelling
6.2.2 Cold | Cold
6.2.3 Dry Heat | DryHeat
6.2.4 Damp heat steady state | DampHeat
6.2.5 AC mains voltage variations | VoltVariations
6.2.6 Variation in voltage supplied by external 12V and 24 V road vehicle batteries | BatteryV
6.3 Tests for disturbances
6.3.1 AC mains voltage dips and short interruptions | VoltDipsInterrupt
6.3.2 Electrical Bursts | BurstsMains
6.3.3 Radiated, radio-frequency, electromagnetic fields | RadiatedRF,EMFields
6.3.4 Conducted, radio-frequency fields | ConductedRF,EMFields
6.3.5 Electrostatic discharge | ElecDischarges
Mechanical shock | MechShock
6.3.6 Storage temperature | StorageTemp
6.3.7 Random vibration | RandVibration
6.4 Tests to assess the submitted calibrations
6.4.1 Accuracy and precision at reference conditions | Accuracy,r&R-Calibration
6.4.2 Sample temperature sensitivity | STS-Calibration

## Authority&Synopsis
1 AUTHORITY RESPONSIBLE FOR THIS REPORT
Organisation name:
Address:
Report number:
Application number:
Testing period: |  | to
Issue date of this Report:
Approver name:
Approver signature:
Stamp(s) if applicable:
2 SYNOPSIS OF THE RESULTS OF THE EXAMINATION AND TESTS
The tested samples of the type fulfils ALL the applicable requirements in OIML R xxx (201x):
Pass/fail
Comments:

## Summary
3 SUMMARY OF THE RESULTS OF THE EXAMINATION AND TESTS
3.1 Examinations
(^) Contains specifications set by the national responsible body in the 'SpecConfirm' sheet.
R xxx Part 1 - Metrological / technical / software requirements |  | Result | Location / ref
3 | ^Units of measurement
4 | Metrological requirements
4.1 | ^Applicable grains and PMB measuring ranges – specification
4.2 | ^Instrument environmental operating temperature – specification
4.3 | ^Grain sample operating temperature – specification
4.4 | ^Influence quantities – specification
4.5 | ^Maximum permissible error and other accuracy requirements
4.7 | Requirements for calibrations |  | 3.2 Performance Tests Summary and test reports
4.8 | Error due to variations in influence quantities
4.9 | Error due to changes in the instrument over time
5 | Technical requirements
5.1 | Checking facilities
5.2 | ^Manufacturer’s manual
5.3 | Markings
5.4 | ^Sample input and calibration selection
5.5 | Instrument construction
5.6 | Level indicating means
5.7 | Presentation of the measured value
6 | Requirements for software-controlled devices and security
6.1 | ^Specification of software requirements
6.2 | ^Data storage
6.3 | ^Data transmission
6.4 | ^Provision for software and calibration security
6.3 | Software documentation
9.4 | ^Maintenance and reconfiguration of the approved software
9.5 | In-field updates to grain calibrations
R xxx Part 2
7.1.2 | Documentation file
3.2 Performance Tests
R xxx Annex C: Type evaluation tests |  | Result | Location
C.4 | Tests for time related effects
C.4.1 | Instrument warm-up time
C.4.2 | Instrument drift and instability
C.5 | Tests for influence variations within the rated operating conditions
C.5.1 | Instrument levelling
C.5.2 | Cold
C.5.3 | Dry heat
C.5.4 | Damp heat
C.5.5 | AC mains voltage variation
C.5.6 | Variation in voltage supplied by external road vehicle batteries
C.6 | Tests for disturbances
C.6.1 | AC mains voltage dips, short interruptions and voltage variations
C.6.2 | Bursts (transients) on AC mains
C.6.3 | Radiated radiofrequency, electromagnetic susceptibility
C.6.4 | Conducted radio-frequency fields
C.6.5 | Electrostatic discharges
C.6.6 | Storage temperature (extreme shipping conditions)
C.6.7 | Random vibration
(continued next page)
R xxx Annex C: Type evaluation tests (continued) |  | Result | Location
C.7 | Assessment of calibrations in the submitted type
GT1 | *
GT2 | *
GT3 | *
GT4 | *
GT*
C.7.1 | Accuracy and precision at reference conditions
GT1 | *
GT2 | *
GT3 | *
GT4 | *
GT*
C.7.1 | Accuracy and precision at reference conditions
GT1 | *
GT2 | *
GT3 | *
GT4 | *
GT*
* Insert Calibration name and version number
Note on J9: 201210 Leida Queddeng (NMIA): Example - Leave "Location / ref" blank. Print report to PDF and then insert respective page number in field and re-printing.

## ApplicationGenInfo
4 GENERAL INFORMATION ABOUT THE APPLICATION
4.1 Manufacturer
Name:
Address:
4.2 Applicant
Organisation:
Contact name:
Address:
Email address:
Phone: |  |  |  | Fax:
Application date: |  |  |  | Ref number:
Applicant is authorised by the manufacturer (documented evidence) | [6 empty] | Yes/ no
Comments:
4.3 Testing laboratories involved in the tests
Complete this form for each test laboratory -
Organisation name:
Address:
Application number:
Tests by this laboratory:
Testing period: |  |  |  |  | to
Name(s) of test engineer(s):
Laboratory accredited by:
Accreditation number: |  |  |  |  | Expiry:
Accreditation includes R xxx: |  |  | Yes/ no |  | Edition:
Details of relevant peer assessement or assessment by other means:
Details, if any tests have been performed at another location than the laboratory premises:
Responsible person - name:
Date signed:
Signature:
Stamp(s) if applicable:
Comments:
4.4 General information concerning type
4.4.1 Description of the instrument (key technical characteristics and intended applications)
e.g. benchtop near infrared (NIR) protein measuring instrument with calibrations for wholegrain wheat and barley
4.4.2 Information indicated on the instrument
Manufacturer trademark:
Year of manufacture:
Type designation:
Model number(s) (if applic):
Electrical power markings:
Software ID (if applic):
Other descriptor/marking:
Comments:
4.4.3 Information on sample units
Serial number |  | Model number |  |  | Manufacture mode (prototype / production) |  | Year
Add additional rows if over three sample units are submitted.
Comments:
4.4.4 Relevant external/internal photographs taken during the examination and tests
4.5 Accessories supplied by the applicant
Batteries (if applicable): |  | Type |  | Vnom |  | Number required
Data printer (if applic):
External data storage (if applic):
Cables:
Other accesories:
4.6 Information on sample instruments
In case the tests and evaluation are valid for more versions, give full details of the types, versions,
measuring ranges, etc.:
Justification for the selection of the sample unit(s):
4.7 Adjustments and modifications
Adjustments, modifications and repairs made to the sample unit(s) during the testing:
4.8 Results of previous tests that were taken into account
Details:
4.9 Additional information concerning type
4.9.1 Instrument limitations of use as declared by the supplier
Calibration dependent
Sample PMB content (%) |  | Min: |  | Max:
Sample moisture content (%) |  | Min: |  | Max:
Operating temperature (°C) |  | Min: |  | Max:
Grain sample temperature (°C) |  | Min: |  | Max:
Maximum Δt (°C) |  | or ΔtC,max |  | & ΔtH,max
If limits are dependent on the grain calibration (i.e. answer is 'Yes') specify values/range in 4.9.2 instead.
Comments:
4.9.2 Information on the submitted calibrations
Calibration principle:
Calibration number |  | GT1 |  |  | GT2 |  | etc.*
Calibration name:
Version number:
Displayed name:
Date submitted:
Calibration limitations of use as declared by the supplier -
Applicable type(s) of grain:
PMB measuring range:
Min/max moisture content:
tC,sample and tH,sample:
Δtmax or ΔtCmax and ΔtHmax:
Regression information -
Approx number of data points:
Data sources, date range:
Reference method(s):
Other validation result:
(e.g. SD, SEP)
Default bias (if applic):
Default slope (if applic):
Other characteristic:
*Copy table into additional pages if more than two calibrations are submitted for examination.
Comments:
4.9.3 Additional information (e.g. connection equipment, interfaces, etc.)
4.10 Documentation supplied by applicant
Date received | Document title and/or reference number |  |  |  | Description (include version number if applicable)
Insert additional rows as required.
4.11 Test equipment and grain used in type evaluation
4.11.1 Test equipment
Instrument/ equipment | Make/ model | Serial # | Parameter applied/ measured |  | Calibrated range(s) | Test(s) used
Insert additional rows as required.
Details e.g. ~ equipment set-up for ESD and EMS tests
~ details of simulations
~ confidence intervals for uncertainty estimations
4.11.2 Grain test samples
Information about the grain reference materials (RMs) used in tests:
Reference method used to generate whole-grain certified reference materials (CRMs):

## SpecConfirm
5 EXAMINATION DETAILS
5.1 Checklist - Specifications to be confirmed by the national authority
Metrological and technical specifications of the national responsible body
3.3^ | The national responsible body has specified the basis moisture content (MB) for expressing the protein content of each grain type. The examiner has entered or attached the national specification.
|  | Grain type name |  | MB
Comments: e.g. ref ID and location in attached publication
use next page if required
4.1 | Applicable grains and PMB measuring ranges – specification
4.1.1^ | The national responsible body has specified commercially important PMB ranges for the grains in R xxx Table 3 on which PMB measurements are subject to national approval. The examiner has entered or attached the national specification for the PMB measurement ranges.
|  | GT name | Min PMB | Max PMB
Comments: e.g. ref ID and location in attached publication
use next page if required
4.2 | Instrument environmental operating temperature – specification
4.2.2^ | The national responsible body has specified the range of ambient temperatures (tC to tH) in which the instrument can be used to take PMB measurements for commercial purposes. The range tC to tH includes 10 ºC to 30 ºC. The examiner has entered the values for tC and tH or attached the national specification.
|  | tC |  | national spec
tH
Comments: e.g. ref ID and location in attached publication
4.2.3(b)^ | The supplier has requested for a wider environmental temperature operating range to be adopted as tC to tH AND the national responsible body has accepted for this particular type evaluation. The examiner has indicated the revised values for tC to tH.
|  | revised tC |  | only for this case
revised tH
Comments:
4.3 | Grain sample operating temperature – specification
4.3.1 | Specification of the sample temperature range
4.3.1.2^ | The national responsible body has specified the range of sample temperatures (tC,sample to tH,sample) in which the instrument can be used to take PMB measurements for commercial purposes. The range tC,sample to tH,sample includes 10 ºC to 30 ºC. The examiner has entered or attached the values of tC,sample and tH,sample for each grain type.
|  | GT name | tC,sample | tH,sample
Comments: e.g. ref ID and location in attached publication
use next page if required
4.3.1.3(b)^ | The supplier has requested for a wider sample temperature range to be adopted as tC,sample to tH,sample AND the national responsible body has accepted for this particular type evaluation. The examiner has entered or attached the revised values of tC,sample and tH,sample for each grain type.
|  | GT name | rev tC,sample | rev tH,sample
Comments: e.g. ref ID and location in attached publication
use next page if required
4.3.2 | Specification of the sample and instrument maximum temperature differential (Δtmax)
4.3.2.2^ | The national responsible body has specified a Δtmax in which the instrument can be used to take commercial PMB measurements. Δtmax (or ΔtC,max and ΔtH,max, if unequal about tref) ≥10 ºC. The examiner has entered or attached the values of Δtmax (or ΔtC,max and ΔtH,max) for each grain type.
|  | GT name | Δtmax or ΔtC,max | & ΔtH,max
Comments: e.g. ref ID and location in attached publication
use next page if required
Additional space for comments, adopted requirements, etc. Reference applicable OIML R xxx clauses.
Continued - Metrological and technical specifications of the national responsible body
4.3.2.3(b)^ | The supplier has requested for a larger maximum differential to be adopted as Δtmax (or ΔtC,max and ΔtH,max) AND the national responsible body has accepted for this type evaluation. The examiner has entered or attached the revised values of Δtmax (or ΔtC,max and ΔtH,max) for each grain type.
|  | GT name | rev Δtmax / rev ΔtC,max | & rev ΔtH,max
Comments:
use next page if required
4.4 | Influence quantities – specification
4.4.1^ | The rated operating ranges for influence factors specified by the national responsible body conform to the default values in the international standard. The examiner has entered or attached any ranges (that differ from OIML R xxx clause 4.4.1) and provided reasons for deviations.
|  | Influence | Revised | Revised
|  | factor | min | max
Comments: e.g. Reason for deviation
use next page if required
4.4.2^ | The disturbances specified by the national responsible body conform to the default values in the international standard. The examiner has entered or attached any settings (that differ from OIML R xxx clause 4.4.2) and provided reasons for deviations.
|  | Disturbance | Revised | Revised
|  | setting | min | max
Comments: e.g. Reason for deviation
use next page if required
4.5 | Maximum permissible error (MPE) and other accuracy requirements
4.5.2^ | The accuracy requirements specified by the national responsible body for each grain type conform to the international standard. The examiner has indicated any MPEs or limits (that differ from OIML R xxx Table 4) and provided reasons for deviations.
|  | GT | Test | Revised
|  | name | parameter | MPE/limit
Comments: e.g. Reason for deviation
use next page if required
4.5.3^ | The MPE for each grain type has been scaled in accordance with the nationally-specified MB and the recommended rounding method has been performed.
Comments:
4.5.4^ | Reference method specified for the national responsible body for PMB measurements is based on an international standard. The examiner has inserted or attached the test procedure and provided reasons for any deviations.
Comments: e.g. ref ID of attached publication
Annex A^ | The national responsible body has specified the relationship between the measured nitrogen content (N) and the protein content (PM or PMB) for each grain type. The examiner has entered or attached the national specification.
|  | GT name | Correlation
Comments: e.g. ref ID and location in attached publication
use next page if required
5.2 | Manufacturer's manual
5.2.3(a)^ | Language(s) other than the official language are accepted by the national responsible body. Examiner has noted all language(s) accepted for the publication.
Comments: Specify language(s).
5.4 | Sample input and calibration selection
5.4.2.2(a)^ | A sample smaller than the international recommendation is permitted by the national responsible body. The examiner has indicated the smallest size permitted by the national responsible body and the reason for deviation from the international standard.
Comments: e.g. Minimum allowable sample size.
5.4.2.3^ | Minimum guidelines for the sampling of cereals for testing - based on international standards - have been specified by the national responsible body. The examiner has referenced or attached the test procedure and provided reasons for deviations from international standard.
Comments: e.g. ref ID of attached publication or reason for deviation.
Additional space for comments, adopted requirements, etc. Reference applicable OIML R xxx clauses.
Software specifications of the national responsible body
6(a)^ | N/A' except if instruments or modules are operated by software. The national responsible body has attempted to adopt the software requirements in OIML R xxx clause 6. The examiner has attached an alternative assessment criteria and report template (where applicable) and provided reasons for deviations from the international recommendation.
Comments:
If the answer to 6(a)^ is 'No' or 'N/A', the following considerations may not apply.
6(b)^ | The possible consequences arising from any fraud or misuse facilitated by the software have been considered by the national responsible body. Severity I is maintained as the appropriate risk level and procedure A methods suffice for validation in most cases. The examiner has provided reasons for deviations from procedure A.
Comments:
6.1 | Specification of the software requirements
6.1.8(a)^ | The national responsible body requires instruments to be equipped with an internal recording element and/or a communication interface that permits interfacing with an external recording element.
Comments:
Data transmission in insecure environment
6.1.9(a)^ | In regards to the trading practices of the country where the type approval is sought- Legally relevant measurement data can be stored outside the measuring instrument or transmitted in an insecure environment before the PMB values are used for commercial purposes.
Comments:
6.1.9(b)^ | N/A' except if the answer to 6.1.9(a)^ is 'Yes'. The national responsible body requires instruments to meet the requirements in OIML R xxx clause 6.3 if it is possible store or transmit the measurement data in an insecure enviroment.
Comments:
6.1.9(c)^ | N/A' except if the answer to 6.1.9(b)^ is 'Yes'. The national responsible body deems the risk associated with instruments that can be used in an insecure environment as severity II and requires data protection via cryptographic means.
Comments:
6.1.9(d)^ | N/A' except if the answer to 6.1.9(c)^ is 'Yes'. The national responsible body requires the effectiveness of data encryption facilities to be validated via Procedure B methods (extended examination).
Comments:
6.2 | Data storage
6.2.1(a)^ | Storage of the legally relevant PMB measurements and data is required by the national responsible body.
Comments:
6.4 | Provision for software and calibration security
6.4.1(a)^ | The national responsible body has adopted all the guidance in Annex G on sealing methods, as the means of ensuring the requirements in clause 6.4.1 are met.
Comments:
6.4.2.4(a)^ | The national responsible body has restricted access to any of the device-specific (i.e. calibration or configuration) parameters. The examiner has entered or attached the details of any access restrictions.
Comments: e.g. Details of any access restrictions to device-specific parameters
9.4 | Maintenance and reconfiguration of the approved software
9.4(a)^ | Software changes on a verified instrument by a 'Traced Update' (instead or in addition to 'Verified Updates') is permitted by the national responsible body.
Comments:
9.5 | In-field updates to grain calibrations
9.5.1(a)^ | The national responsible body allows for updates to PMB calibrations to accommodate for seasonal variations.
9.5.1(b)^ | The national responsible body requires retention of data collected during the current and/or recent years that are used for calibration adjustments.
Comments:
9.5.3^ | A calibration update process which is not a Verified Update, is permitted by the national responsible body, provided that the instrument fulfils the requirements for Traced Updates at type evaluation.
Comments:
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
N/A
Yes
No
Yes
No
N/A
Yes
No
N/A
Yes
No
N/A
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No

## ExamDetails
5 EXAMINATION DETAILS
5.2 Examination checklist - metrological and technical requirements
(^) Clause is associated with specifications set by the national responsible body in the 'SpecConfirm' sheet.
Checks on the requirements within R xxx Part 1 |  |  | Findings
3 | Units of measurement
3.1 | The measurement of the protein content in a grain sample is expressed in percentage by mass (% w/w). The percentage symbol alone (%) is also permitted.
Comments:
3.2 | For each type of grain, the measured protein content is expressed at one basis moisture content (MB). The scaling of the protein content at the actual “as is” moisture content (PM) to the protein content at the basis moisture content (PMB) is in accordance with Equation 1.
Comments:
4 | Metrological requirements
4.1 | Applicable grains and PMB measuring ranges – specification
4.1.2 | The grain types and the corresponding PMB measuring range that can be analysed by the submitted type of instrument has been declared by the supplier and inserted in section 4.9.1 or 4.9.2 of the 'ApplicationGenInfo' sheet. The PMB measuring ranges includes the range specified by the national responsible body for the grain type.
Comments:
4.2 | Instrument environmental operating temperature – specification
4.2.1 | N/A' except if answers to 4.2.2^ and 4.2.3(b)^ are 'No' (i.e. a limits for the environmental temperature has NOT been specified by the national responsible body). Performance test(s) demonstrate that the submitted type provides accurate PMB measurements in all the environmental temperatures possible in the country/region.
Comments:
4.2.3(a) | A rated operating range for the environmental temperature has been declared by the supplier as shown in section 4.9.1 or 4.9.2 of the 'ApplicationGenInfo' sheet. Unless a wider range is specified in 4.2.3(b)^, the supplier range is equivalent to the national specification in 4.2.2^.
Comments:
4.3 | Grain sample operating temperature – specification
4.3.1 | Specification of the sample temperature range
4.3.1.1 | N/A' except if answers to 4.3.1.2^ and 4.3.1.3(b)^ are 'No', i.e. a limited range for the sample temperature (tC,sample to tH,sample) has NOT been specified by the national responsible body. Performance test(s) demonstrate that the submitted type provides accurate PMB measurements on samples that are between 2 ºC to 45 ºC in temperature.
Comments:
4.3.1.3(a) | A limited range the sample temperature has been declared by the supplier as shown in section 4.9.1 or 4.9.2 of the 'ApplicationGenInfo' sheet. Unless a wider range is specified in 4.3.1.3(b)^, the supplier range(s) are equivalent to the national specification in 4.3.1.2^.
Comments:
4.3.2 | Specification of the sample and instrument maximum temperature differential (Δtmax)
4.3.2.1 | NA' except if answers to 4.3.2.2^ and 4.3.2.3(b)^ are 'No', i.e. Δtmax (maximum limit for the sample and instrument temperature differential when that latter is a tref) has NOT been specified by the national responsible body. Performance test(s) demonstrate that the instrument at reference temperature (tref), provides accurate PMB measurements on samples within the range tC,sample to tH,sample regardless of the magnitude of Δt.
Comments:
(^) Clause is associated with specifications set by the national responsible body in the 'SpecConfirm' sheet.
Continued - Checks on the requirements within R xxx Part 1 |  |  | Findings
4.3.2.3(a) | A maximum limit for Δt has been declared by the supplier as shown in section 4.9.1 or 4.9.2 of the 'ApplicationGenInfo'. Unless a larger Δtmax is specified in 4.3.2.3(b)^, supplier limit(s) for Δt are equivalent to the national specification in 4.3.2.2^.
Comments:
4.3.3 | Provisions in absence of a manufacturer-specified sample temperature range
4.3.3(a) | N/A' except if answer to 4.3.1.3(a) is 'No'. Performance test(s) demonstrate that the submitted type provides accurate PMB measurements. Where relevant, the examiner has attached national provisions (e.g. operating procedures for ground grain protein analysers) aimed to mitigate risks associated with the lack of sample temperature detection/ compensation facilities.
Comments: e.g. ref ID of attached publication
4.3.3(b) | N/A' except if answer to 4.3.2.3(a) is 'No'. Performance test(s) demonstrate that the submitted type provides accurate PMB measurements. Where relevant, the examiner has attached national provisions (e.g. operating procedures for ground grain protein analysers) aimed to mitigate risks associated with the lack of Δt detection/ compensation facilities.
Comments: e.g. ref ID of attached publication
5 | Technical requirements
5.1 | Checking facilities
5.1.1 | Suppression of PMB measured values following a significant fault. There is provision for the instrument to automatically and clearly indicate when a significant fault has occurred by an appropriate error message, unambiguous warning or blanking the display.
Comments:
5.1.2 | Suppression of PMB measured values outside of operating ranges
5.1.2.1(a) | There is provision for the instrument to automatically and clearly indicate if the following type-approved operating range is exceeded, by an appropriate error message, unambiguous warning or blanking the display: Range tC to tH specified in 4.2.2^ or 4.2.3(b)^.
Comments:
5.1.2.1(b) | There is provision for the instrument to automatically and clearly indicate if the following type-approved operating range(s) is exceeded, by an appropriate error message, unambiguous warning or blanking the display: Range(s) tC,sample to tH,sample specified in 4.3.1.2^ or 4.3.1.3(b)^.
Comments:
5.1.2.1{c) | There is provision for the instrument to automatically and clearly indicate if the following type-approved operating limit(s) are exceeded, by an appropriate error message, unambiguous warning or blanking the display: Δtmax for each grain type as specified in 4.3.2.2^ or 4.3.2.3(b)^.
Comments:
5.1.2.2 | The instrument shall automatically prevent further measurements as long as the respective influence factor or sample characteristic remains outside the type-approved ranges.
Comments:
5.1.2.3 | The operator is not required to judge the precise ambient temperature and the temperature of the sample in order to make an accurate measurement.
Comments:
5.1.3 | Suppression of PMB values or warnings outside of the approved measuring range. There is provision for the instrument to automatically suppress PMB measured values that are outside of the type-approved measurement range for the calibration, unless it is accompanied by an appropriate error message or unambiguous warning.
Comments:
5.1.4 | N/A' for instruments which do not require any warm up time. After the instrument is switched on, there is provision for any measured values to be suppressed until the operating temperature necessary for accurate measurement has been attained .
Comments:
(^) Clause is associated with specifications set by the national responsible body in the 'SpecConfirm' sheet.
Continued - Checks on the requirements within R xxx Part 1 |  |  | Findings
5.2 | Manufacturer's manual
5.2.1 | There is a manual to be provided with each protein measuring instrument, that describes the installation, operation, and routine maintenance of the instrument and its accessories.
Comments:
5.2.2(a) | The manual includes the name and address of the manufacturer.
Comments:
5.2.2(b) | The manual includes type of the instrument with which it is intended to be used.
Comments:
5.2.2{c) | The manual includes the date of issue.
Comments:
5.2.2(d) | The manual includes the types of grain for which the instrument is designed to be used.
Comments:
5.2.2{e) | The manual includes the limitations of use of the instrument.
Comments:
5.2.3(b) | N/A' except if answer to 5.2.3(a)^ is 'Yes'. The user/owner manual is supplied in all the accepted or official language(s). The translations appear to be accurate.
Comments:
5.3 | Markings
5.3.1(a) | The instrument is clearly and permanently marked with the manufacturer’s name or mark.
Comments:
5.3.1(b) | Instrument is clearly and permanently marked with the model designation.
Comments:
5.3.1{c) | Instrument is clearly and permanently marked with the serial number given by the manufacturer.
Comments:
5.3.1(d) | Provision is made for application of a type approval mark.
Comments:
5.3.2 | Markings have been grouped together in a clearly visible location, either on a permanently attached nameplate or on part of the instrument. The required information is readily observable without disassembly.
Comments:
5.3.3 | All operational controls, indications, indicating switches, features, light displays and push button shall be clearly identifiable. Keys visible only to the operator need only be marked to the extent that a trained operator can understand the function of each key.
Comments:
5.4 | Sample input and calibration selection
5.4.1 | Selection of calibration on the instrument
5.4.1.1 | On instruments that have a different calibration for each grain type, the user is able to select the calibration applicable for the sample to be analysed.
Comments:
5.4.1.2 | The selection of the calibration on the user interface is unambiguous and visible to all parties present, i.e. during the measurement it can be verified that the selected (displayed) calibration corresponds with the sample analysed.
Comments:
5.4.2 | Sampling and minimum sample size
5.4.2.1 | The operator is not required to judge the precise volume or weight required by the instrument to make an accurate PMB measurement.
Comments:
5.4.2.2(b) | NA' except if the answer to 5.4.2.2(a)^ is 'No'. The size of the sample analysed by the instrument is at least 100 g or 400 kernels or seeds (which ever is smaller).
Comments:
5.5 | Instrument construction
5.5.1 | Nothing observed in the design and construction of the instrument and accessory equipment should make it prone to inaccuracy, malfunction and fraud under normal service conditions.
Comments:
5.5.2 | Day to day forces on the parts of the instruments shall not compromise the accuracy of measurements.
Comments:
5.5.3 | The instrument housing protects the main components from dust and moisture.
Comments:
(^) Clause is associated with specifications set by the national responsible body in the 'SpecConfirm' sheet.
Continued - Checks on the requirements within R xxx Part 1 |  |  | Findings
5.5.4 | NA' except if the instrument analyses ground/milled samples-. The manufacturer of the protein measuring instrument has designated the type(s) of mill to be used and included a unit with the submission so its suitability for the measurement process be assessed during type evaluation. The examiner has indicated the appropriate mill type(s).
Comments: e.g. Specification of mill type.
5.6 | Level indicating means
5.6.1(a) | Based on manufacturer specifications and preliminary tests, tilting the instrument in any upright direction by up to 5% (approximately 3º) may reduce the accuracy of the instrument.
Comments:
5.6.1(b) | N/A' except if the previous answer is 'Yes'. The instrument is equipped with a level indicator and level adjustment means to reduce the likelihoood of being tilted.
Comments:
5.6.2 | N/A' except if the answer to 5.6.1(a) is 'Yes'. The level indicating means is readable without any instrument disassembly.
Comments:
5.7 | Presentation of the measured value
5.7.1 | The instrument is equipped with a digital indicating element which does not display any protein content values before the end of the measurement cycle.
Comments:
5.7.2 | Measurement results are displayed as percent protein by mass (%) the MB. Subdivisions of this unit are in terms of decimal subdivisions (not fractions).
Comments:
5.7.3(a) | The display on the type allows the protein content value to be determined with a resolution of at least 0.1% PMB.
Comments:
5.7.3(b) | Sample instruments submitted for type evaluation permit 0.01% PMB resolution.
Comments:
5.7.4(a) | The type is multi-constituent measuring instrument (e.g. it measures grain moisture content in addition to protein content)
Comments:
5.7.4(b) | NA' except if the previous answer is 'Yes'. Appropriate labels are displayed or recorded to make it clear which constituent is associated the displayed or recorded measured values.
Comments:
5.7.5 | The height for the digits used to display protein content is at least 10 mm. Numbers and symbols of units are presented in accordance with OIML D 2.
Comments:
5.3 Examination checklist - requirements for software-controlled devices and security
NOTE: All the following are not applicable 'N/A' except if the answer to 6(a)^ is 'Yes'
(^) Clause is associated with specifications set by the national responsible body in the 'SpecConfirm' sheet.
Checks on the software requirements within R xxx Part 1 |  |  | Findings
6.1 | Specification of the software requirements
6.1.1 | For instruments and modules operated by software, the manufacturer has described or declared how the software is implemented within the instrument or module. Examiner to indicate whether the software is embedded or on an universal computer system.
| Details of level I validation method: | AD
Insert details or the reference relevant sections of an attached software validation report.
6.1.2 | The legally relevant software is clearly identifiable via a unique software version or a checksum. Examiner to indicate whether the software version or the checksum is displayed or printed out on command during operational mode, or displayed during the start-up procedure.
| Details of level I validation method(s): | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
6.1.3 | The legally relevant measuring algorithms and functions are appropriate and functionally correct. Examiner to conduct further examinations and metrological tests if necessary.
| Details of level I validation method(s): | AD + VFTM
Insert details or the reference relevant sections of an attached software validation report.
6.1.4(a) | Optional assessment during type evaluation- Conformity of the legally relevant software (in the submitted units) to that in the approved type has been verified at level (b) described in D 31 clause 5.2.5. Examiner to confirm the identity of parts of the legally relevant source code, and for the rest of the software, the identity of the legally relevant functions described in the documentation.
| Details of level I validation method(s): | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
6.1.4(b) | Selected functions or parts of the source code can be modified.
Comments:
6.1.4(c) | N/A' except if the previous answer is 'Yes'. It is possible to detect software variations, e.g. via checksum values.
| Details of level I validation method(s): | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
6.1.5 | There is provision to make further measurements impossible when a significant fault is detected.
| Details of level I validation method(s): | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
6.1.6(a) | The software of the instrument is separated into a legally-relevant part and non-relevant parts.
Comments:
6.1.6(b) | N/A' except if the previous answer is 'Yes'. The requirements of D 31 clause 5.2.1.2 for Separation of software parts have been fulfilled.
| Details of level I validation method: | AD
Insert details or the reference relevant sections of an attached software validation report.
6.1.7(a) | The instrument uses an internal or external universal computer.
Comments:
6.1.7(b) | Cryptographic data protection is implemented in the software of the instrument.
Comments:
6.1.7(c) | N/A' except if the answer to 6.1.7(a) and/or 6.1.7(b) is 'Yes'. The legally relevant software can be operated only in the environment specified for its correct functioning. If necessary to secure the correct functioning of the legally relevant software, the operating system has been fixed to a defined invariant configuration.
| Details of level I validation method(s): | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
6.1.8(b) | N/A' except if the answer to 6.1.8(a)^ is 'Yes'. The instrument is equipped with the required recording element(s).
Comments:
(^) Clause is associated with specifications set by the national responsible body in the 'SpecConfirm' sheet.
Checks on the software requirements within R xxx Part 1 |  |  | Findings
6.1.8{c) | NA' except if answers to 6.1.8(a)^ and 6.1.8(b) are 'Yes'. Correspondence between the displayed information and remote recording element has been verified during performance test(s).
Comments:
6.2 | Data storage All 'N/A' except if the answer to 6.2.1(a)^ is 'Yes'
| Details of level I validation method for 6.2: | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
6.2.1(b) | The measurement data is stored automatically when the measurement is finished.
Comments:
No protein content values are recorded before the end of the measurement cycle.
Comments:
The storage device has sufficient permanency to ensure that data are not corrupted.
Comments:
There is sufficient memory for storage of the required measurement data to be used at a later time.
Comments:
6.2.2 | The measurement value stored is accompanied by all relevant information necessary for future legally relevant use.
Comments:
The measurement records include: test sample identifier, measurement date and time, unique identification of the instrument, grain type, PMB results and units as displayed, calibration version ID, error messages and constituent labels (on multi-constituent meters).
Comments:
6.3 | Data transmission All 'N/A' except if the answer to 6.1.9(b)^ is 'Yes'
| Details of level I validation method: | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
6.3.1 | The data is protected by software means to guarantee the authenticity and integrity.
Comments:
There is provision for the data to be discarded or marked unusable if an irregularity is detected.
Comments:
6.3.3 | There is provision so that the measurement is not inadmissibly influenced by a transmission delay.
Comments:
6.3.4 | There is provision to ensure measurement data is not lost if a transmission interruption occurs because the network services become unavailable (e.g. the measurement process should stop).
Comments:
6.3.2 | NA' except if the answers to both 6.1.9(c)^ and 6.1.9(d)^ are 'Yes'. Data is protected by cryptographic means. A seal is broken if a confidential key is input or read.
| Details of level II validation method: | AD + VFTSw + SMT
Insert details or the reference relevant sections of an attached software validation report.
6.4 | Provision for software and calibration security
6.4.1 | Sealing
6.4.1(b) | N/A' except if unrestricted or remote access to sealable parameters is possible AND the answer to 6.4.1(a)^ is 'Yes'. Annex G Checklist-
An event logger has been provided which includes the following: an event counter, the parameter ID, the date and time of change, and the new value of the parameter changed. For changes to multiple calibration constants (e.g. 10 or more), the calibration version number may be presented rather than the calibration constants.
Comments:
The event logger automatically retains the ID of the parameter changed, the date and time of change, and the new value of the parameter.
Comments:
If the instrument does not include a printer, there is provision for attachment of a printer, which can print the contents of the audit trail upon demand.
Comments:
The event logger has the capacity to retain records equal to 25 times the number of sealable parameters in the instrument. Not more than 1000 records are required.
Comments:
(^) Clause is associated with specifications set by the national responsible body in the 'SpecConfirm' sheet.
Checks on the software requirements within R xxx Part 1 |  |  | Findings
6.4.1(b) | Non-sealable parameters that are routinely changed in normal operation, cannot be accessed in the mode which allows sealable parameters to be adjusted.
cont. |  |  | Comments:
The event counter cannot be reset by the operator and has the capacity of at least 1000 values (e.g. 000 to 999).
Comments:
The event counter increments appropriately. In the case of an event event logger, if an adjustment mode is entered but no changes are made, this does not constituite an event and the counter must not increment.
Comments:
The event logger drops the oldest event when the memory capacity is full and a new entry is saved.
Comments:
The audit trail information is capable of being retained in memory for at least 30 days while the instrument is without power.
Comments:
Accessing the audit train information for review is separate from the mode used to enter or modify sealable parameters.
Comments:
Audit trail information can be accessed without requiring the removal of any parts other than the normal requirements to inspect the integrity of a physical seal or to use a key (for a panel lock).
Comments:
The printed audit trail is readily interpretable by an enforcement official.
Comments:
Event logger information is printed in order from the most recent event to the oldest event.
Comments:
If the event logger information is printed on more than one line per event, information is printed in blocks which are readily understandable.
Comments:
6.4.1(c) | N/A' except if unrestricted or remote access to sealable parameters is possible AND the answer to 6.4.1(a)^ is 'No'. Provision has been made for appropriate sealing by mechanical, electronic and/or cryptographic means, making any change that affects the metrological integrity of the instrument impossible or evident. Calibrations, zero-setting and test point adjustments are sealed.
| Details of level I validation method: | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
6.4.2 | Safeguards against fraudulent use
| Details of level I validation method: | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
6.4.2.1 | The legally relevant software has been secured against unauthorised modification, loading or changes by swapping of the memory device. If the instrument has an operating system or an option to load software, additional means to mechanical sealing have been considered.
Comments:
6.4.2.2 | Only clearly documented functions are allowed to be activated by the user interface, which have been realised in such a way that it does not facilitate fraudulent use.
Comments:
6.4.2.3 | Parameters that fix the legally relevant characteristics of the measuring instrument shall be secured against unauthorised modification. The current parameter settings can be printed or displayed on demand.
Comments:
6.4.2.4(b) | NA' except if the answer to 6.4.2.4(a)^ is 'Yes'. The instrument complies with any access restrictions to adjustable parameters.
Comments:
9.4 | Maintenance and reconfiguration of the approved software
9.4(b) | N/A' except if the answer to 9.4(a)^ is 'Yes'. Implementation of the software fulfils the requirements for Traced Updates in D 31 clause 5.2.6.2.
| Details of level I validation method: | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
Checks on software requirements within R xxx Parts 1 & 2 |  |  | Findings
9.5 | In-field updates to grain calibrations
9.5.2 | Calibration version
The calibration constants that are adjustable and unique calibration names, or calibration version numbers can be displayed and printed on demand.
Comments:
9.5.3 | Security of calibrations and reverification
| Details of level I validation method: | AD + VFTSw
Insert details or the reference relevant sections of an attached software validation report.
There is provision for only authorised persons to change grain calibrations.
Comments:
Changes to the grain calibrations of the instrument are impossible unless recorded in an audit trail.
Comments:
There is provision for an error message to be automatically displayed if calibration constants are electronically corrupted and to make further measurements impossible.
Comments:
6.5 | Software documentation
The submitted software documentation is complete. Examiner has indicated the inclusion or exclusion of the following information with the submission-
Comments:
Description of the legally relevant software, incorporating how the requirements are met
Comments:
Description of the operating system security
Comments:
Description of the software sealing method(s)
Comments:
Overview of the system hardware, highlighting any hardware components that are deemed legally relevant or performing legally relevant functions
Comments:
Description of the accuracy of the algorithms
Comments:
Declaration of the hardware and software environment, including minimum resources and configuration necessary for correct functioning of the instrument
Comments:
Description of the user interface, menus and dialogues
Comments:
Description of the software identification which has to be clearly assigned to the legally relevant functions
Comments:
Clear instructions on how to check the actual software identification against the reference number as listed in the type approval certificate
Comments:
List of commands of each hardware interface of the measuring instrument/ electronic device/ sub-assembly
Comments:
List of durability errors that are detected by the software
Comments:
Description of data sets stored or transmitted
Comments:
List of significant faults that are detected and a description of the detecting algorithm
Comments:
Operating manual which clearly identifies all operational controls, indications, and features
Comments:
5.4 Examination checklist - documentation requirements
Checks on the requirements within R xxx Part 2 |  |  | Findings
7.1.2 | Documentation file
The submitted documentation file is complete. Examiner has indicated the inclusion or exclusion of the following information with the submission-
Comments:
General description of the instrument
Comments:
General characteristics of the instrument
Comments:
Drawings of general arrangement and details of metrological interest
Comments:
Description of calibrations submitted for approval
| Comments: |  | Comments:
Descriptions and characteristic data for all devices and sub-assemblies of the instrument
Comments:
Technical description, drawings and plans of devices, sub-assemblies
Comments:
Declarations of the manufacturer
| Comments: |  | Comments:
Samples of all intended print-outs
| Comments: |  | Comments:
Information concerning special cases
| Comments: |  | Comments:
Results of tests performed by the manufacturer using protocols from Parts 2 and 3
Comments:
Certificates of other type approvals or separate tests
| Comments: |  | Comments:
Drawing or photo of the instrument
| Comments: |  | Comments:
Manufacturer’s manual according to OIML R xxx clause 5.2. NOTE: Ensure consistency with the responses in 5.2.1, 5.2.2 and 5.2.3(b).
Comments:
Software documentation according to OIML R xxx clause 6.3
| Comments: |  | Comments:
Other evidence to support the assumption that the design and characteristics of the type comply with the requirements of R xxx
Comments:
Pass
Fail
Pass
Fail
Pass
Fail
Yes
No
Pass
Fail
N/A
Yes
No
Pass
Fail
N/A
Yes
No
Pass
Fail
N/A
Pass
Fail
N/A
Pass
Fail
N/A
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
N/A
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
N/A
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
N/A
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
N/A
Yes
No
Pass
Fail
N/A
Pass
Fail
N/A
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Yes
No
Pass
Fail
N/A
Pass
Fail
Pass
Fail
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Yes
No
Pass
Fail
N/A
Yes
No
Yes
No
Pass
Fail
N/A
Pass
Fail
N/A
Pass
Fail
N/A
Pass
Fail
N/A
Yes
No
Pass
Fail
N/A
N/A
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
N/A
N/A
Pass
Fail
N/A
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
N/A
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail
N/A
Pass
Fail
N/A
Pass
Fail
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Pass
Fail
Pass
Fail
Pass
Fail
Pass
Fail

## WarmUp
6 TYPE EVALUATION TESTS
6.1 Tests for time-related effects
6.1.1 Instrument warm-up time [ref. OIML R xxx Annex C clause C.4.1]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
|  | Not warm | Warm
Ambient temp (t): |  |  |  | °C
Ambient RH: |  |  |  | %
Date commenced: |  |  |  | ddmmyyyy
Time commenced: |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  | Nominal values (%)
name
GT | Min | Max | ID |  | PMB | M
Instrument | Mean PMB |  | Error shift |  | RESULTS SUMMARY (pass / fail)
| Not warm | Warm | Warm |  | Comments |  | Instrument | Type
1
2
Raw data entry
| PMB (Not warmed-up) |  |  |  | PMB (Warm-up time elapsed)
| Instr. 1 | Instr. 2 | Comments |  | Instr. 1 | Instr. 2 | Comments
1
2
3
4
5

## DriftInstability
6.1.2 Instrument drift and instability [ref. OIML R xxx Annex C clause C.4.2]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
7.1.2 Instrument 2 ID:
|  | Before testing received | After 4-6 wks
Ambient temp (t): |  |  |  | °C
Ambient RH: |  |  |  | %
Date commenced: |  |  |  | ddmmyyyy
Time commenced: |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  |  | Nominal values (%)
name
GT | Min | Max | number | ID |  | PMB | M
1
2
3
Sample number | Instrument | Mean PMB |  | Error shift |  | RESULTS SUMMARY (pass / fail)
|  | Start | 4 - 6 wks | After 4 - 6 weeks |  | Comments |  | Instrument | Type
1 | 1
2
3
1 | 2
2
3
Raw data entry
Sample number | PMB (Before testing, except warm-up test) |  |  |  | PMB (After 4 - 6 weeks)
| Instr. 1 | Instr. 2 | Comments |  | Instr. 1 | Instr. 2 | Comments
1
2
3

## Levelling
6.2 Tests for influence factor variations within the rated operating ranges
6.2.1 Instrument levelling [ref. OIML R xxx Annex C clause C.5.1]
Observer: | [6 empty] | Displayed | Error shift limits |  | Sample info |  | Nominal values (%)
Type/ application #: | [6 empty] | name
Instrument 1 ID: | [6 empty] | GT | Min | Max | ID |  | PMB | M
Instrument 2 ID:
|  | Start - ref | Influence |  |  | Recovery |  |  | General comments on test (e.g. tilt orientation):
|  | Level | 1st tilt | 2nd tilt | * | Level
Tilt angle: | [6 empty] | ° (angluar degrees)
Ambient temp (t): | [6 empty] | °C
Ambient RH: | [6 empty] | %
Date commenced: | [6 empty] | ddmmyyyy
Time commenced: | [6 empty] | hh:mm
* Add other tilt orientations if required
Instrument | Mean PMB |  |  |  |  | Error shift (at tilt orientation) |  |  |  | RESULTS SUMMARY (pass / fail)
| Start - ref | 1st | 2nd | * | Recovery | 1st | 2nd | * | Recovery | Comments |  | Instrument | Type
1
2
Raw data entry
PMB (Start -ref) |  |  | PMB (1st orientation) |  |  | PMB (2nd orientation) |  |  | PMB (*) |  |  | PMB (Recovery)
Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments
* Add other tilt orientations if required

## Cold
6.2.2 Cold [ref. OIML R xxx Annex C clause C.5.1]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
Spare instrument ID:
|  | Starf ref | Cold (tC) | Recovery
EUT ambient temp: |  |  |  |  | °C
EUT ambient RH: |  |  |  |  | %
Sample temp: |  |  |  |  | °C
Spare ambient temp: |  |  |  |  | °C
Spare ambient RH: |  |  |  |  | %
Date commenced: |  |  |  |  | ddmmyyyy
Time commenced: |  |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  |  | Nominal values (%) |  | Sample stability | Correct for recov
name
GT | Min | Max | number | ID |  | PMB | M | Δ PMB | Yes/ No
1
2
3
Sample | Instrument | Error shift |  |  | RESULTS SUMMARY (pass/fail)
number |  | Cold (tC) | Recovery | Corrected | Comments |  | Instrument | Type
1 | 1
2
3
1 | 2
2
3
Sample | Mean PMB (Start - ref) |  |  | Mean PMB (Cold) |  | Mean PMB (Recovery)
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Sample | PMB (Start - ref) |  |  | PMB (Cold) |  | PMB (Recovery) |  |  | Comments
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Comments

## DryHeat
6.2.3 Dry heat [ref. OIML R xxx Annex C clause C.5.3]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
Spare instrument ID:
|  | Starf ref | Dry H (tH) | Recovery
EUT ambient temp: |  |  |  |  | °C
EUT ambient RH: |  |  |  |  | %
Sample temp: |  |  |  |  | °C
Spare ambient temp: |  |  |  |  | °C
Spare ambient RH: |  |  |  |  | %
Date commenced: |  |  |  |  | ddmmyyyy
Time commenced: |  |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  |  | Nominal values (%) |  | Sample stability | Correct for recov
name
GT | Min | Max | number | ID |  | PMB | M | Δ PMB | Yes/ No
1
2
3
Sample | Instrument | Error shift |  |  | RESULTS SUMMARY (pass/fail)
number |  | Dry H (tH) | Recovery | Corrected | Comments |  | Instrument | Type
1 | 1
2
3
1 | 2
2
3
Sample | Mean PMB (Start - ref) |  |  | Mean PMB (Dry Heat) |  | Mean PMB (Recovery)
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Sample | PMB (Start - ref) |  |  | PMB (Dry Heat) |  | PMB (Recovery) |  |  | Comments
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Comments

## DampHeat
6.2.4 Damp heat [ref. OIML R xxx Annex C clause C.5.4]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
7.2.4 Instrument 2 ID:
Spare instrument ID:
|  | Starf ref | Damp H (tH) | Recovery
EUT ambient temp: |  |  |  |  | °C
EUT ambient RH: |  |  |  |  | %
Sample temp: |  |  |  |  | °C
Spare ambient temp: |  |  |  |  | °C
Spare ambient RH: |  |  |  |  | %
Date commenced: |  |  |  |  | ddmmyyyy
Time commenced: |  |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  |  | Nominal values (%) |  | Sample stability | Correct for recov
name
GT | Min | Max | number | ID |  | PMB | M | Δ PMB | Yes/ No
1
2
3
Sample | Instrument | Error shift |  |  | RESULTS SUMMARY (pass/fail)
number |  | Damp H (tH) | Recovery | Corrected | Comments |  | Instrument | Type
1 | 1
2
3
1 | 2
2
3
Raw data entry - Damp heat [ref. OIML R xxx Annex C clause C.5.4]
Sample | Mean PMB (Start - ref) |  |  | Mean PMB (Damp H) |  | Mean PMB (Recovery)
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Sample | PMB (Start - ref) |  |  | PMB (Damp heat) |  | PMB (Recovery) |  |  | Comments
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Comments

## VoltVariations
6.2.5 AC mains voltage variations [ref. OIML R xxx Annex C clause C.5.5]
Observer: | [6 empty] | Displayed | Error shift limits |  | Standard deviation | Sample info |  | Nominal values (%)
Type/ application #: | [6 empty] | name
Instrument 1 ID: | [6 empty] | GT | Min | Max | Max | ID |  | PMB | M
Instrument 2 ID: | [9 empty] | 0.1
|  | Start - ref | Influence |  | Recovery |  |  | General comments on test:
|  | Unom | Unom +10% | Unom -15% | Unom
Test voltage: | [5 empty] | V
Test frequency: | [5 empty] | Hz
Ambient temp (t): | [5 empty] | °C
Ambient RH: | [5 empty] | %
Date commenced: | [5 empty] | ddmmyy
Time commenced: | [5 empty] | hh:mm
Instrument | Standard deviation replicate PMB |  |  |  | Mean PMB |  |  |  | Error shift |  |  | RESULTS SUMMARY (pass / fail)
| Start - ref | Unom +10% | Unom -15% | Recovery | Start - ref | Unom +10% | Unom -15% | Recovery | Unom +10% | Unom -15% | Recovery | Comments | Instrument | Type
1
2
Raw data entry
# | PMB (Start -ref) |  |  | PMB (Unom +10%) |  |  | PMB (Unom -15%) |  |  | PMB (Recovery)
| Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments
1
2
3
4
5
6
7
8
9
10

## BatteryV
6.2.6 Variation in voltage supplied by external 12V and 24 V road vehicle batteries [ref. OIML R xxx Annex C clause C.5.6]
Observer: | [6 empty] | Displayed | Error shift limits |  | Sample info |  | Nominal values (%)
Type/ application #: | [6 empty] | name
Instrument 1 ID: | [6 empty] | GT | Min | Max | ID |  | PMB | M
Instrument 2 ID:
|  | Start - ref | Influence |  | Recovery |  |  | General comments on test:
|  | Unom | Low V | High V | Unom
Test voltage: | [5 empty] | V
Test frequency: | [5 empty] | Hz
Ambient temp (t): | [5 empty] | °C
Ambient RH: | [5 empty] | %
Date commenced: | [5 empty] | ddmmyy
Time commenced: | [5 empty] | hh:mm
Instrument | Mean PMB |  |  |  | Error shift |  |  | RESULTS SUMMARY (pass / fail)
| Start - ref | Low V | High V | Recovery | Low V | High V | Recovery | Comments | Instrument | Type
1
2
Raw data entry
# | PMB (Start -ref) |  |  | PMB (Low V) |  |  | PMB (High V) |  |  | PMB (Recovery)
| Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments
1
2
3
4
5
6
7
8
9
10

## VoltDipsInterrupt
7.3 Tests for disturbances
7.3.1 AC mains voltage dips and short interruptions [ref. OIML R xxx Annex C clause C.6.1]
Observer: | [6 empty] | General comments on test:
Type/ application #:
Instrument ID:
Displayed | Fault limits |  | Sample info | Nominal values (%)
name
GT | Min | Max | ID | PMB | M
Test (min voltage indicated)
|  | Start testing | A | B | C | D | End testing
|  |  | 0 × Unom | 0 × Unom | 0.7 × Unom | 0 × Unom
Test voltage: |  |  | 0 | 0 | fill cell C12 | 0 |  | V
Min voltage duration: |  |  | 0.5 | 1 | 25/30^ | 250/300^ |  | cycles
Test frequency: | [7 empty] | Hz
Ambient temp (t): | [7 empty] | °C
Ambient RH: | [7 empty] | %
Date: | [7 empty] | ddmmyy
Time: | [7 empty] | hh:mm
^These values are for 50 Hz and 60 Hz respectively
All faults within limits (yes / no) |  | All insignificant# (yes / no) |  | RESULT (pass / fail)
|  |  |  | Type | Comments
#A fault exceeding the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Raw data entry
| Ref PMB | A |  | B |  | C |  | D
|  | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)

## RadiatedRF,EMFields
6.3.3 Radiated, radio-frequency, electromagnetic fields [ref. OIML R xxx Annex C clause C.6.3]
Observer: | [5 empty] | Displayed | Fault limits
Type/ application #: | [5 empty] | name
Instrument ID: | [5 empty] | GT | Min | Max
General comments on test:
All faults within limits (yes / no) |  | All insignificant# (yes / no) |  | RESULT (pass / fail)
|  |  |  | Type | Comments
#A fault exceeding the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Disturbance settings (indicate values if alternative settings are applied)
Frequency range: |  |  | MHz |  | Frequency range: |  | 26* - 2000 | MHz
Modulation: |  |  |  |  | Modulation: |  | 80% AM, 1 kHz sine wave
Field strength: |  |  | V/ m |  | Field strength: |  | 10 | V/ m
Step size: |  |  |  |  | Step size: |  | 1 | %
Rate of sweep: |  |  |  |  | Rate of sweep: |  | 0.0015 | decade/ s
*Testing from 80 MHz is permitted.
Antenna:
Circular polarisation (yes / no):
Note: Two orthogonal polarisations shall be tested if the answer is 'No'.
Example radiated EMS testing scheme:
| Antenna polarisation: | 1 - Vertical |  |  |  | 2 - Horizontal
| Facing EUT side: | Front | Left | Right | Back | Front | Left | Right | Back
Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp (t):
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp (t):
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp (t):
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp (t):
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp (t):
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp (t):
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp (t):
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp (t):
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.

## BurstsMains
6.3.2 Bursts (transients) on AC mains [ref. OIML R xxx Annex C clause C.6.2]
Observer: | [10 empty] | Start test | End test
Type/ application #: | [9 empty] | Test voltage: |  |  | V
Instrument ID: | [9 empty] | Test frequency: |  |  | Hz
[10 empty] | Ambient temp (t): |  |  | °C
General comments on test: | [9 empty] | Ambient RH: |  |  | %
[10 empty] | Date: |  |  | ddmmyy
[10 empty] | Time: |  |  | hh:mm
Mean ref PMB:
1
Displayed | Fault limits |  | Sample info | Nominal values (%) | [5 empty] | 2
name | [9 empty] | 3
GT | Min | Max | ID | PMB | M |  |  |  |  | 4
5
All faults within limits (yes / no) |  | All insignificant# (yes / no) |  | RESULT (pass / fail)
|  |  |  | Type | Comments
#A fault exceeding the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Disturbance severity |  |  | Amplitude: |  | 1 | kV
|  |  | Repetition rate: |  | 5 | kHz
Raw data entry - Bursts (transients) on AC mains [ref. OIML R xxx Annex C clause C.6.2]
Connection |  | L → G |  |  |  | N → G |  |  |  | PE → G
Polarity |  | ( + ) |  | ( - ) |  | ( + ) |  | ( - ) |  | ( + ) |  | ( - )
|  | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)

## ConductedRF,EMFields
6.3.4 Conducted, radio-frequency, electromagnetic fields [ref. OIML R xxx Annex C clause C.6.5]
Observer: | [9 empty] | Start test | End test
Type/ application #: | [8 empty] | Test voltage: |  |  | V
Instrument ID: | [8 empty] | Test frequency: |  |  | Hz
[9 empty] | Ambient temp (t): |  |  | °C
Displayed | Fault limits |  | Sample info | Nominal values (%) |  |  |  |  | Ambient RH: |  |  | %
name | [8 empty] | Date: |  |  | ddmmyy
GT | Min | Max | ID | PMB | M |  |  |  | Time: |  |  | hh:mm
Mean ref PMB:
1
General comments on test: | [8 empty] | 2
3
4
5
All faults within limit (yes / no) |  | All insignificant# (yes / no) |  | RESULT (pass / fail)
|  |  |  | Type | Comments
#A fault exceeding the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Disturbance settings (indicate values if alternative settings are applied)
Frequency range: |  |  | MHz |  |  |  | Frequency range: |  | 0.15 - 80 | MHz
Modulation: | [6 empty] | Modulation: |  | 80% AM, 1 kHz sine wave
RF amplitude (50 Ω): |  |  | V (e.m.f.) |  |  |  | RF amplitude (50 Ω): |  | 10 | V (e.m.f.)
Raw data entry - Radiated, radio-frequency, electromagnetic fields [ref. OIML R xxx Annex C clause C.6.3]
Cable or interface: | [5 empty] | Cable or interface:
Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
* | [5 empty] | *
Add additional rows if required.
Cable or interface: | [5 empty] | Cable or interface:
Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments | Frequency (MHz) | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
* | [5 empty] | *
Add additional rows if required.

## ElecDischarges
6.3.5 Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Observer: | [10 empty] | Start test | End test
Type/ application #: | [9 empty] | Test voltage: |  |  | V
Instrument ID: | [9 empty] | Test frequency: |  |  | Hz
[10 empty] | Ambient temp (t): |  |  | °C
Displayed | Fault limits |  | Sample info | Nominal values (%) | [5 empty] | Ambient RH: |  |  | %
name | [9 empty] | Date: |  |  | ddmmyy
GT | Min | Max | ID | PMB | M |  |  |  |  | Time: |  |  | hh:mm
Mean ref PMB:
1
General comments on test: | [9 empty] | 2
3
4
5
All faults within limits (yes / no) |  | All insignificant# (yes / no) |  | RESULT (pass / fail)
|  |  |  | Type | Comments
#A fault exceeding the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Disturbance settings
Direct discharge mode (contact, paint penetration or air) | [5 empty] | Refer to diagram of coupling plane positions for applying discharges indirectly.
2 kV | [5 empty] | Notes: H = horizontal, V = vertical
4 kV
6 kV | [8 empty] | V2
8 kV
[7 empty] | V1 |  | EUT |  | V3
front
V4
Polarity of indirect discharges (positive / negative):
It is permitted to only test one at polarity.
IEC 61000-4-2 specifies to select the most sensitive polarity.
Raw data entry - Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Direct discharge application
|  | 2 |  |  | 4 |  |  | 6 |  |  | 8
Voltage (kV) |  | PMB |  | Comments | PMB |  | Comments | PMB |  | Comments | PMB |  | Comments
Polarity |  | ( + ) | ( - ) |  | ( + ) | ( - ) |  | ( + ) | ( - ) |  | ( + ) | ( - )
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)
Continued - Raw data entry - Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Indirect discharge application
|  | V1 |  |  |  | V2 |  |  |  | V3
Voltage (kV) |  | PMB |  |  | Comments | PMB |  |  | Comments | PMB |  |  | Comments
Polarity |  | 2 | 4 | 6 |  | 2 | 4 | 6 |  | 2 | 4 | 6
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)
Continued - Raw data entry - Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Continued - Indirect discharge application
Voltage (kV) |  | V3 |  |  |  | H
|  | PMB |  |  | Comments | PMB |  |  | Comments
Polarity |  | 2 | 4 | 6 |  | 2 | 4 | 6
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)

## StorageTemp
6.3.6 Storage temperature (extreme shipping conditions) [ref. OIML R xxx Annex C clause C.6.6]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
|  | Start test testing | End test testing
Ambient temp: |  |  |  | °C
Ambient RH: |  |  |  | %
Date: |  |  |  | ddmmyy
Time: |  |  |  | hh:mm
Displayed | Fault limits |  | Sample info | Nominal values (%)
name
GT | Min | Max | ID | PMB | M
Instrument | All faults within limit (Yes / no) |  | All insignificant# (Yes / no) |  | RESULTS SUMMARY (Pass / fail)
[5 empty] | Comments |  | Instrument | Type
1
2
#A fault exceeding the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Disturbance severity |  |  | Minimum temp: |  | -20 | °C
|  |  | Maximum temp: |  | 55 | °C
Raw data entry
BEFORE disturbance
# | PMB
| Instr. 1 | Instr. 2
1
2
3
4
5
6
7
8
9
10
Mean PMB
AFTER disturbance
| Instrument 1 |  |  |  |  | Instrument 2
| PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) |  | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no)
1 |  |  |  |  | 1
2 |  |  |  |  | 2
3 |  |  |  |  | 3
4 |  |  |  |  | 4
5 |  |  |  |  | 5
6 |  |  |  |  | 6
7 |  |  |  |  | 7
8 |  |  |  |  | 8
9 |  |  |  |  | 9
10 |  |  |  |  | 10
Comments |  |  |  |  | Comments

## RandVibration
6.3.7 Random vibrations [ref. OIML R xxx Annex C clause C.6.7]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument ID:
|  | Start test | End test
Ambient temp (t): |  |  |  | °C
Ambient RH: |  |  |  | %
Date: |  |  |  | ddmmyy
Time: |  |  |  | hh:mm
Displayed | Fault limits |  | Sample info | Nominal values (%)
name
GT | Min | Max | ID | PMB | M
All faults within limits (yes / no) |  | All insignificant (yes / no) |  | RESULT (pass / fail)
|  |  |  | Type | Comments
#A fault exceeding the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Disturbance severity |  |  | Total frequency range: |  |  | 10 - 150 | Hz
|  |  | Total RMS level: |  |  | 7 | ms-2
|  |  | ASD level 10 Hz – 20 Hz: |  |  | 1 | m2s-3
|  |  | ASD level 20 Hz – 150 Hz |  |  | -3 | dB/octave
|  |  | Number of directions: |  |  | 3
|  |  | Duration per axis: |  |  | 60 | minutes
Raw data entry
|  | PMB BEFORE |  | AFTER disturbance
|  |  |  | PMB | Fault | Within limit (yes / no) | Insignificant (yes / no) | Comments
| 1 |  | 1
| 2 |  | 2
| 3 |  | 3
| 4 |  | 4
| 5 |  | 5
| Mean PMB |  | 6
7
8
9
10

## Accuracy,r&R-Calibration
6.4 Tests to assess the submitted calibrations
6.4.1 Accuracy and precision [ref. OIML R xxx Annex C clause C.7.1]
Observer: | [5 empty] | General comments on test settings:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
|  | GT1 |  | GT2 |  | GT3 |  | GT4 |  | GT*
|  | Start ref | End ref | Start ref | End ref | Start ref | End ref | Start ref | End ref | Start ref | End ref
Ambient temp (t): | [11 empty] | °C
Ambient RH: | [11 empty] | %
Date commenced: | [11 empty] | ddmmyyyy
Time commenced: | [11 empty] | hh:mm
Submitted calibrations | [5 empty] | RESULTS SUMMARY (pass / fail)
Number | Calibration name |  | Version | Disp. name | MB (%) | Accuracy | Repeat. | Reprod. | Comments
GT1
GT2
GT3
GT4
GT*
Limits
Calibration | y(bar) |  | SEP | Pooled SD | SDDI
number | Min | Max | Max | Max | Max
GT1
GT2
GT3
GT4
GT*
* Add cells for additional calibrations if necessary
Calibration | Sample info |  | Reference values |  | Sample info |  | Reference values |  | Sample info |  | Reference values
number | number | ID | PMB | M | number | ID | PMB | M | number | ID | PMB | M
GT1 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
GT2 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
* Add cells for additional samples if necessary
Calibration | Sample info |  | Reference values |  | Sample info |  | Reference values |  | Sample info |  | Reference values
number | number | ID | PMB | M | number | ID | PMB | M | number | ID | PMB | M
GT3 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
GT4 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
GT*
* Add cells for additional calibrations and/or samples if necessary
[9 empty] | GT1 | on Instr.1 | GT1 | on Instr. 2
The requirement for y(bar) was fulfilled and all functions operated as designed (pass or fail):
Calibration | Sample | Mean PMB |  |  | Error (yi) |  | Pooled yi, i.e. y(bar)
number | number | reference | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MinLimit | MaxLimit
GT1 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
Add cells for additional samples if necessary
[9 empty] | GT2 | on Instr.1 | GT2 | on Instr. 2
The requirement for y(bar) was fulfilled and all functions operated as designed (pass or fail):
Calibration | Sample | Mean PMB |  |  | Error (yi) |  | Pooled yi, i.e. y(bar)
number | number | reference | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MinLimit | MaxLimit
GT2 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
GT*
* Add cells for additional calibrations and/or samples if necessary
[9 empty] | GT1 | on Instr.1 | GT1 | on Instr. 2
The requirement for SEP was fulfilled and all functions operated as designed (pass or fail):
Calibration | Sample | Reference | PMB (j=1) |  | Error (y,j=1) |  | SEP
number | number | PMB | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MaxLimit
GT1 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
* Add cells for additional samples if necessary
[9 empty] | GT2 | on Instr.1 | GT2 | on Instr. 2
The requirement for SEP was fulfilled and all functions operated as designed (pass or fail):
Calibration | Sample | Reference | PMB (j=1) |  | Error (y,j=1) |  | SEP
number | number | PMB | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MaxLimit
GT2 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
GT*
* Add cells for additional calibrations and/or samples if necessary
[9 empty] | GT1 | on Instr.1 | GT1 | on Instr. 2
The requirement for repeatability was fulfilled and all functions operated as designed (pass or fail):
Calibration | Sample | SD |  | SD2 (variance) |  | Pooled SD
number | number | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MaxLimit
GT1 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
* Add cells for additional samples if necessary
[9 empty] | GT2 | on Instr.1 | GT2 | on Instr. 2
The requirement for repeatability was fulfilled and all functions operated as designed (pass or fail):
Calibration | Sample | SD |  | SD2 (variance) |  | Pooled SD
number | number | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MaxLimit
GT2 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
GT*
* Add cells for additional calibrations and/or samples if necessary
[9 empty] | GT1 | on Instr.1 | GT1 | on Instr. 2
The requirement for reproducibility was fulfilled and all functions operated as designed (pass or fail):
Calibration | Sample | Mean PMB |  | d | SDDI | SSDI
number | number | Instr. 1 | Instr. 2 |  |  | MaxLimit
GT1 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
* Add cells for additional samples if necessary
[9 empty] | GT2 | on Instr.1 | GT2 | on Instr. 2
The requirement for reproducibility was fulfilled and all functions operated as designed (pass or fail):
Calibration | Sample | Mean PMB |  | d | SDDI | SSDI
number | number | Instr. 1 | Instr. 2 |  |  | MaxLimit
GT2 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
GT*
* Add cells for additional calibrations and/or samples if necessary
Raw data entry - Accuracy and precision [ref. OIML R xxx Annex C clause C.7.1]
Calibration | Sample | PMB @ ref |  | Comments | Sample | PMB @ ref |  | Comments | Sample | PMB @ ref |  | Comments
number | number | Instr. 1 | Instr.2 |  | number | Instr. 1 | Instr.2 |  | number | Instr. 1 | Instr.2
GT1 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
* Add cells for additional samples if necessary
Continued - Raw data entry - Accuracy and precision [ref. OIML R xxx Annex C clause C.7.1]
Calibration | Sample | PMB @ ref |  | Comments | Sample | PMB @ ref |  | Comments | Sample | PMB @ ref |  | Comments
number | number | Instr. 1 | Instr.2 |  | number | Instr. 1 | Instr.2 |  | number | Instr. 1 | Instr.2
GT2 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
GT*
* Add cells for additional calibrations and/or samples if necessary
Note on G260: Leida Queddeng (NMIA): SQRT(sum variances / n) (min n = 30)
Note on H260: Leida Queddeng (NMIA): SQRT(sum variances / n) (min n = 30)
Note on G296: Leida Queddeng (NMIA): SQRT(sum variances / n) (min n = 30)
Note on H296: Leida Queddeng (NMIA): SQRT(sum variances / n) (min n = 30)

## STS-Calibration
6.4.2 Sample temperature sensitivity (STS) [OIML Rxxx-2, clause C.7.2]
Observer: | [8 empty] | Start ref | tref - ΔtC | Recov 1 | tref + ΔtH | Recov 2
Type/ application #: | [6 empty] | Ambient temp: | [6 empty] | °C
Instrument 1 ID: | [6 empty] | Ambient RH: | [6 empty] | %
Instrument 2 ID: | [6 empty] | Sample temp: | [6 empty] | °C
[7 empty] | Date commenced: | [6 empty] | ddmmyyyy
[7 empty] | Time commenced: | [6 empty] | hh:mm
General comments on test settings:
Limits
Calibrations submitted and tested | [6 empty] | Pooled error shift |  | RESULTS SUMMARY
number | Calibration name |  | Version | Displayed name |  | MB (%) | Min | Max | pass / fail | Comments
GT1
GT2
GT3
GT4
Calibration | Sample info |  |  | Nomimal values (%) |  |  |  | Calibration | Sample info |  |  | Nomimal values (%)
number | moisture | number | ID | PMB | M |  |  | number | moisture | number | ID | PMB | M
GT1 | low m | 1 | [5 empty] | GT3 | low m | 1
|  | 2 | [7 empty] | 2
|  | 3 | [7 empty] | 3
| high m | 4 | [6 empty] | high m | 4
|  | 5 | [7 empty] | 5
|  | 6 | [7 empty] | 6
GT2 | high m | 1 | [5 empty] | GT4 | high m | 1
|  | 2 | [7 empty] | 2
|  | 3 | [7 empty] | 3
| low m | 4 | [6 empty] | low m | 4
|  | 5 | [7 empty] | 5
|  | 6 | [7 empty] | 6
Instrument 1 only
Calibration | Sample info |  | Mean PMB |  |  |  |  | Error shift |  | Pooled error shift |  | RESULT | Limits pooled e/shift
number | moisture | number | Start ref | tref - ΔtC | Recov 1 | tref + ΔtH | Recov 2 | tref - ΔtC | tref + ΔtH | tref - ΔtC | tref + ΔtH | pass / fail | Min | Max
GT1 | low m | 1
2
3
| high m | 4
5
6
GT2 | high m | 1
2
3
| low m | 4
5
6
GT3 | low m | 1
2
3
| high m | 4
5
6
GT4 | high m | 1
2
3
| low m | 4
5
6
[8 empty] | All functions operated as intended: |  |  |  |  | pass / fail
Instrument 2 only
Calibration | Sample info |  | Mean PMB |  |  |  |  | Error shift |  | Pooled error shift |  | RESULT | Limits pooled e/shift
number | moisture | number | Start ref | tref - ΔtC | Recov 1 | tref + ΔtH | Recov 2 | tref - ΔtC | tref + ΔtH | tref - ΔtC | tref + ΔtH | pass / fail | Min | Max
GT1 | low m | 1
2
3
| high m | 4
5
6
GT2 | high m | 1
2
3
| low m | 4
5
6
GT3 | low m | 1
2
3
| high m | 4
5
6
GT4 | high m | 1
2
3
| low m | 4
5
6
[8 empty] | All functions operated as intended: |  |  |  |  | pass / fail
Raw data entry - Sample temperature sensitivity (STS) [ref. OIML R xxx Annex C clause C.7.2]
Calibration | Sample info |  | PMB (start ref) |  | PMB (tref - ΔtC) |  | PMB (recov 1) |  | PMB (tref + ΔtH) |  | PMB (recov 2) |  | Comments e.g. functionality
number | moisture | number | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2
GT1 | low m | 1
2
3
| high m | 4
5
6
GT2 | high m | 1
2
3
| low m | 4
5
6
*
* Add cells for additional calibrations if necessary
Continued - Raw data entry - Sample temperature sensitivity (STS) [ref. OIML R xxx Annex C clause C.7.2]
Calibration | Sample info |  | PMB (start ref) |  | PMB (tref - ΔtC) |  | PMB (recov 1) |  | PMB (tref + ΔtH) |  | PMB (recov 2) |  | Comments e.g. functionality
number | moisture | number | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2
GT3 | low m | 1
2
3
| high m | 4
5
6
GT4 | high m | 1
2
3
| low m | 4
5
6

## FormSelections
NAME | PassOrFail | PassFailNA | YesOrNo | YesNoNA
| pass | pass | yes | yes
| fail | fail | no | no
|  | NA |  | NA